Hardness tester HL200

## Configuration:

|  | No. | Item | Quantity | Remarks |
| --- | --- | --- | --- | --- |
| Standard Configuration | 1 | Main body | 1 |  |
|  | 2 | D type impact device | 1 | With cable |
|  | 3 | Standard test block | 1 |  |
|  | 4 | Cleaning brush (I) | 1 |  |
|  | 5 | Small support ring | 1 |  |
|  | 6 | Alkaline battery | 4 | AA size |
|  | 7 | Manual | 1 |  |
|  | 8 | Instrument case | 1 |  |
|  | 9 | DataPro software for HL200 | 1 |  |
|  | 10 | Communication cable | 1 |  |
|  | 11 | Tool for battery cover | 1 |  |
|  | 12 | Belt | 1 |  |
| Optional Configuration | 11 | Cleaning brush (II) | 1 | For use with other type of impact device |
|  | 12 | Other type of impact devices | 7 | Refer to pictures blew |
|  | 13 | Tool for impact ball | 1 |  |
|  | 14 | Min-printer | 1 |  |
|  | 15 | Other type of support rings | 1 | Refer to pictures blew |
|  | 16 | Other type of impact ball / body | 1 |  |

## Technical Specifications:

- Error and repeatability of displayed value

| No. | Type of impact device | Hardness value of hardness block | Error of displayed value | Repeatability |
| --- | --- | --- | --- | --- |
| 1 | D | 760±30HLD 530±40HLD | ±6 HLD ±10 HLD | 6 HLD 10 HLD |
| 2 | D | 50±5 HLC 30±6 HLC | ±1 HRC ±2 HRC | 1 HRC 2 HRC |
| 3 | D | 700±30HB 500±40HB | ±5 HB ±8 HB | 5 HB 8 HB |

- Measuring range：

(170-960)HLD,(17-68.5)HRC,(19-651)HB,(80-976)HV,(30-100)HS,(59-85)HRA,(13-100)HRB

- Measuring direction：360°（↓↖↗↘↙→←↑）

- Hardness Scale：HL、HB、HRB、HRC、HRA、HV、HS

- Display：dot matrix LCD，128×64 dots

- Data memory：max. 500 groups（relative to impact times 32～1）

- Working voltage：3V（2 AA size alkaline battery）

- Continuous working period：about 50 hours（With backlight off）

- Communication interface：USB1.1

Main Application

- Die cavity of molds

- Bearings and other parts

- Failure analysis of pressure vessel, steam generator and other equipment

- Heavy work piece

- The installed machinery and permanently assembled parts.

- Testing surface of a small hollow space

- Material identification in the warehouse of metallic materials

- Rapid testing in large range and multi-measuring areas for large-scale work piece

## Testing range:

Other type of impact devices

| Material | Method | Impact device |
| --- | --- | --- |
|  |  | D/DC | D+15 | C | G | E | DL |
| Steel and cast steel | HRC | 20～68.5 | 19.3～67.9 | 20.0～69.5 |  | 22.4～70.7 | 20.6～68.2 |
|  | HRB | 38.4～99.6 |  |  | 47.7～99.9 |  | 37.0～99.9 |
|  | HRA | 59.1～85.8 |  |  |  | 61.7～88.0 |  |
|  | HB | 127～651 | 80～638 | 80～683 | 90～646 | 83～663 | 81～646 |
|  | HV | 83～976 | 80～937 | 80～996 |  | 84～1042 | 80～950 |
|  | HS | 32.2～99.5 | 33.3～99.3 | 31.8～102.1 |  | 35.8～102.6 | 30.6～96.8 |
| Cold work tool steel | HRC | 20.4～67.1 | 19.8～68.2 | 20.7～68.2 |  | 22.6～70.2 |  |
|  | HV | 80～898 | 80～935 | 100～941 |  | 82～1009 |  |
| Stainless steel | HRB | 46.5～101.7 |  |  |  |  |  |
|  | HB | 85～655 |  |  |  |  |  |
|  | HV | 85～802 |  |  |  |  |  |
| Grey cast iron | HRC |  |  |  |  |  |  |
|  | HB | 93～334 |  |  | 92～326 |  |  |
|  | HV |  |  |  |  |  |  |
| Nodular cast iron | HRC |  |  |  |  |  |  |
|  | HB | 131～387 |  |  | 127～364 |  |  |
|  | HV |  |  |  |  |  |  |
| Cast aluminium alloys | HB | 19～164 |  | 23～210 | 32～168 |  |  |
|  | HRB | 23.8～84.6 |  | 22.7～85.0 | 23.8～85.5 |  |  |
| BRASS(copper-zinc alloys) | HB | 40～173 |  |  |  |  |  |
|  | HRB | 13.5～95.3 |  |  |  |  |  |
| BRONZE(copper-aluminium/tin alloys) | HB | 60～290 |  |  |  |  |  |
| Wrought copper alloys | HB | 45～315 |  |  |  |  |  |
| Available type of impact device |  | DC: Test hole or hollow cylindrical | D+15: Test groove or reentrant surface | C:Test small, light, thin parts and surface of hardened layer | G: Test large, thick, heavy and rough surface steel | E: Test super high hardness material | DL: Test slender narrow groove or hole |

## Testing conditions:

| Type of impact device | DC(D)/DL | D+15 | C | G | E |
| --- | --- | --- | --- | --- | --- |
| Impacting energy Mass of impact body | 11mJ 5.5g/7.2g | 11mJ 7.8g | 2.7mJ 3.0g | 90mJ 20.0g | 11mJ 5.5g |
| Test tip hardness: Dia. Test tip: Material of test tip: | 1600HV 3mm Tungsten carbide | 1600HV 3mm Tungsten carbide | 1600HV 3mm Tungsten carbide | 1600HV 5mm Tungsten carbide | 5000HV 3mm synthetic diamond |
| Impact device diameter: Impact device length: Impact device weight: | 20mm 86(147)/ 75mm 50g | 20mm 162mm 80g | 20mm 141mm 75g | 30mm 254mm 250g | 20mm 155mm 80g |
| Max. hardness of sample | 940HV | 940HV | 1000HV | 650HB | 1200HV |
| Mean roughness value of sample surface Ra: | 1.6μm | 1.6μm | 0.4μm | 6.3μm | 1.6μm |
| Min. weight of sample: Measure directly Need support firmly Need coupling tightly | >5kg 2～5kg 0.05～2kg | >5kg 2～5kg 0.05～2kg | >1.5kg 0.5～1.5kg 0.02～0.5kg | >15kg 5～15kg 0.5～5kg | >5kg 2～5kg 0.05～2kg |
| Min. thickness of sample Coupling tightly Min. layer thickness for surface hardening | 5mm ≥0.8mm | 5mm ≥0.8mm | 1mm ≥0.2mm | 10mm ≥1.2mm | 5mm ≥0.8mm |
| Size of tip indentation |
| Hardness 300HV | Indentation diameter Depth of indentation | 0.54mm 24μm | 0.54mm 24μm | 0.38mm 12μm | 1.03mm 53μm | 0.54mm 24μm |
| Hardness 600HV | Indentation diameter Depth of indentation | 0.54mm 17μm | 0.54mm 17μm | 0.32mm 8μm | 0.90mm 41μm | 0.54mm 17μm |
| Hardness 800HV | Indentation diameter Depth of indentation | 0.35mm 10μm | 0.35mm 10μm | 0.35mm 7μm | -- -- | 0.35mm 10μm |

## Support rings for Shaped Materials:

Other type of support rings

| No. | Type | Sketch of non-conventional Supporting ring | Remarks |
| --- | --- | --- | --- |
| 1 | Z10-15 |  | For testing cylindrical outside surface R10～R15 |
| 2 | Z14.5-30 |  | For testing cylindrical outside surface R14.5～R30 |
| 3 | Z25-50 |  | For testing cylindrical outside surface R25～R50 |
| 4 | HZ11-13 |  | For testing cylindrical inside surface R11～R13 |
| 5 | HZ12.5-17 |  | For testing cylindrical inside surface R12.5～R17 |
| 6 | HZ16.5-30 |  | For testing cylindrical inside surface R16.5～R30 |
| 7 | K10-15 |  | For testing spherical outside surface SR10～SR15 |
| 8 | K14.5-30 |  | For testing spherical outside surface SR14.5～SR30 |
| 9 | HK11-13 |  | For testing spherical inside surface SR11～SR13 |
| 10 | HK12.5-17 |  | For testing spherical inside surface SR12.5～SR17 |
| 11 | HK16.5-30 |  | For testing spherical inside surface SR16.5～SR30 |
| 12 | UN |  | For testing cylindrical outside surface, radius adjustable R10～∞ |

# Text boxes

- Compact metal case，suitable for use under poor working conditions. Test at any angle, even upside down.

- Wide measuring range. It can measure the hardness of all metallic materials. Direct display of hardness scales HRB, HRC, HV, HB, HS, HL

- Large screen(128×64 dot matrix LCD), showing all functions and parameters. With EL background light.

- Large capacity memory could store 500 groups information.

- Datapro Software to connect with PC via USB port.

- Software calibration function.

- Outline :132*76.2mm    Weight：345g